## Sheet1
140126 InstEvalTemplate.xls
140126 InstEval Stanford LSRIIA2.xlsx
Template form for information on each instrument in the study
There should be one sheet per instrument, all in identical layout.
Instrument Name
LSRII
Instrument Location
Stanford
Folder Name
| SSFF_LSRII |  | NOT FL CHANNELS: | Time | FSC-A | FSC-W | FSC-H | SSC-A | SSC-W | SSC-H
Accommodate this?
| Dye or Sample name | APC | APC-Cy7 | APC-H7 | FITC | PE | PE-Cy7 | PerCP | PerCP-Cy55 | V450 | V500-C | Duke 1011 | 197B | LED+Lasers | Sph8 AE-01 | TF6 FC3MV | LED Series | Sph8 series
| SU2B ABD unit values from Ming | 223145 | 175084 | 175205 | 151918 | 86901 | 194182 | 59257 | 155773 | 264475 | 172187
| Channel name on this instrument ($PnN) | APC-A | APC-Cy7-A | APC-Cy7-A | FITC-A | PE-A | PE-Cy7-A | PerCP-Cy5-5-A | PerCP-Cy5-5-A | Pacific Blue-A | Aqua Amine-A
| Folder name | SU-2B | SU-2B | SU-2B | SU-2B | SU-2B | SU-2B | SU-2B | SU-2B | SU-2B | SU-2B | OtherTests | OtherTests | OtherTests | OtherTests | OtherTests | LED Series | Sph8 series
| File Name | 933723.fcs | 933725.fcs | 933727.fcs | 933729.fcs | 933731.fcs | 933733.fcs | 933735.fcs | 933737.fcs | 933739.fcs | 933741.fcs | 933749.fcs | 933743.fcs | 935319.fcs | 933745.fcs | 933747.fcs | 19 files | 8 files
| Scatter gate 1 channel name ($PnN) | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A | FSC-A |  | FSC-A | FSC-A
| Scatter gate 2 channel name ($PnN) | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A | SSC-A |  | SSC-A | SSC-A
| Number of peaks per file | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 1 | 1 | 1 | 8 | 6 | 1 | 1
| Measurement type (Height, Area) | Area | Area | Area | Area | Area | Area | Area | Area | Area | Area
| Laser wavelength | 640 nm | 640 nm | 640 nm | 488 nm | 532 nm | 532 nm | 488 nm | 488 nm | 405 nm | 405 nm
| Laser power | 40 mW | 40 mW | 40 mW | 50 mW | 150 mW | 150 mW | 50 mW | 50 mW | 50 mW | 50 mW
| Emission filter designation | 670/30 | 780/60 | 780/60 | 525/50+532notch | 575/25 | 780/60 | 682/33 | 710/50 | 450/50 | 515/20
if needed | Detector pass low side |  |  |  | 505 nm -18notch
if needed | Detector pass high side |  |  |  | 550 nm
fcs file evaluations | Capture bead pos Mean
Capture bead pos SD
Capture bead neg Mean
Capture bead neg SD
Duke 1011 Mean
Duke 1011 SD
197B Mean
197B SD
LED fits | Spe/signal unit
Spe/signal unit - StdErr
Bead fits or 1011 | Background SD
Background SD - Std Err
| Dye | APC | APC-Cy7 | APC-H7 | FITC | PE | PE-Cy7 | PerCP | PerCP-Cy55 | V450 | V500-C
Results for comparisons | Spe/ABD unit
Spe/ABD unit - Std Err
Background SD in Spe
Background SD in ABD units
Background SD in ABD units - StdErr
Spe added for 2+2 SD separation
ABD added for 2+2 SD separation
197B Mean in ABD units
LED and 6,8Pk Bead data
Lasers off? YES | LED 01 Mean
LED 01 SD
LED 01 Count
LED 02 Mean
LED 02 SD
LED 02 Count
LED 03 Mean
LED 03 SD
LED 03 Count
LED 04 Mean
LED 04 SD
LED 04 Count
LED 05 Mean
LED 05 SD
LED 05 Count
LED 06 Mean
LED 06 SD
LED 06 Count
LED 07 Mean
LED 07 SD
LED 07 Count
LED 08 Mean
LED 08 SD
LED 08 Count
LED 09 Mean
LED 09 SD
LED 09 Count
LED 10 Mean
LED 10 SD
LED 10 Count
LED 11 Mean
LED 11 SD
LED 12 Mean
LED 12 SD
LED 12 Count
LED 13 Mean
LED 13 SD
LED 13 Count
LED 14 Mean
LED 14 SD
LED 14 Count
LED 15 Mean
LED 15 SD
LED 15 Count
LED 16 Mean
LED 16 SD
LED 16 Count
LED 17 Mean
LED 17 SD
LED 17 Count
LED 18 Mean
LED 18 SD
LED 18 Count
LED 19 Mean
LED 19 SD
LED 19 Count
LED 20 Mean
LED 20 SD
LED 20 Count
LED no signal run? YES | LED no signal Mean
LED no signal SD
LED lasers on run? YES | LED no sig las on Mean
LED no sig las on SD
Sph8 AE01 | Peak 1 Mean
Peak 1 SD
Peak 1 Count
Peak 2 Mean
Peak 2 SD
Peak 2 Count
Peak 3 Mean
Peak 3 SD
Peak 3 Count
Peak 4 Mean
Peak 4 SD
Peak 4 Count
Peak 5 Mean
Peak 5 SD
Peak 5 Count
Peak 6 Mean
Peak 6 SD
Peak 6 Count
Peak 7 Mean
Peak 7 SD
Peak 7 Count
Peak 8 Mean
Peak 8 SD
Peak 8 Count
TF6 FC3MV | Peak 1 Mean
Peak 1 SD
Peak 1 Count
Peak 2 Mean
Peak 2 SD
Peak 2 Count
Peak 3 Mean
Peak 3 SD
Peak 3 Count
Peak 4 Mean
Peak 4 SD
Peak 4 Count
Peak 5 Mean
Peak 5 SD
Peak 5 Count
Peak 6 Mean
Peak 6 SD
Peak 6 Count
Quadratic fit results | Dye | APC | APC-Cy7 | APC-H7 | FITC | PE | PE-Cy7 | PerCP | PerCP-Cy55 | V450 | V500-C
LED | c0
c0 StdErr
c1
c1 StdErr
c2
c2 StdErr
Sph8 AE01 | c0
c0 StdErr
c1
c1 StdErr
c2
c2 StdErr
TF6 FC3MV | c0
c0 StdErr
c1
c1 StdErr
c2
c2 StdErr
Spe estimator RelStdErr
LED | c1 Relative StdErr (%)
Sph8 AE01 | c1 Relative StdErr (%)
TF6 FC3MV | c1 Relative StdErr (%)